Human-Machine Interface, H.M.I. enabled by

Epidermal Electronics System, E.E.S.

ECE445 Spring 2013

Design Review

Woosik Lee, Ohjin Kwon, Nithin Reddy

T.A. : Lydia Majure

February 28, 2013

Table of Contents

Introduction

- Statement of Purpose                                                    (3)

- Benefits & Feature                                                      (3)

Design

  - Block Diagrams                                                        (5)

  - Block Descriptions                                                      (6)

    - Overall Summary                                                 (6)

    - E.E.S.                                                          (7)

    - S.C.U.                                                          (8)

    - Convertor (A/D)                                                 (11)

    - Micro-Controller                                                (12)

  - Schematic of Overall System                                             (13)

  - Simulations & Calculation                                               (15)

Requirements & Verification                                                  (18)

Tolerance Analysis                                                           (20)

Ethical Issues                                                               (21)

Safety                                                                      (22)

Cost & Schedule

A. Cost Analysis                                                          (23)

B. Schedule                                                             (27)

Reference                                                                  (28)

INTROUDUCTION

- Statement of Purpose

The purpose of this project is to design an epidermal electronics system (E.E.S.) and apply this concept for human-machine interface (H.M.I.). The micro/nanotechnology development has access to many types of motion sensors such as Nintendo Wii game controller. However, users still need to hold the controller physically with their hands in order to control any machine or video game. This project would allow those people with neurological and muscular disorders to control machines without having to physically access them. Moreover, we can use it for handling critical risk materials such as scientific studying field. It would make it easier to conduct EMG tests without the hassle of all the bulky equipment. Here, our team wishes to address this issue by directly mounting the electronic system onto skin epidermis, such that human motions can be directly translated into the controlling system.

- Benefits & Feature.

Our project will be able to provide accurate EMG readings when compared to formal EMG equipment readings and will have an added bonus of being extremely mobile. There is no question of discomfort as the E.E.S. which consists of stretchable, meandering structures that naturally follow the soft, curvilinear skin morphology. This allows us to avoid using conductive gels which can cause irritation and other side effects on our skin.

The E.E.S. on the skin gives high-quality and highly accurate electrophysiological signals for an extremely long time (a week) without the use of adhesives and/or tapes because of its capability to be laminated, conform ally, onto the skin. The E.E.S. has some other individual features that help the portability:

- Ultrathin

- Lightweight

- Low modulus (E.E.S. that has the matching mechanics to the skin epidermis).

For the E.E.S., Signal Conditioning Unit (S.C.U) and easily programmable microcontroller are included in our design. Our goal is to control quad-rotor using EMG signals measured on forearm by E.E.S.

Design

- Block Diagrams

Fig. 1. Overall Diagram

Fig. 2. S.C.U. Diagram

Fig. 3. Flowchart of microcontroller program

- Block Descriptions

B.1. Overall Summary

- Epidermal Electronic System (E.E.S) : The EMG sensor will include three electrodes, each in a form of a filamentary serpentine (FS) mesh with exposed metal that contacts the skin directly, for measurement (MEA), ground (GND), and reference (REF).

- Signal Conditioning Unit (S.C.U.) : This unit will receive signals from EES, detect certain amplitude, and amplify the signal. This Unit must contain band pass filter, gain amplifier, and noise filter system.

- Convertor (A/D) : It is essential for the received signal to be converted to digital in order for it to be analyzed. This unit will convert the received analog signal (bio-potentials) from the skin and convert it to a digital signal that is then be analyzed by the microcontroller.

- Micro-Controller : The regulation of the received signal will be done with the microcontroller. It will compare the received signal and with the assistance of predefined programming the necessary action is brought about.

- Computer : The commands will go through the computer in order to control machine wirelessly. For now, we are planning to use quad-rotor as our machine.

- Power Supply : The power supply will provide power to the various components of the circuit. We will use a battery to

B.2. E.E.S. (Epidermal Electronic System)

The one of the principle technology of electrodes is used in order to measure electrophysiological processes in the muscle (electromyograms, EMGs). To minimize impedances between the electrodes and the human body and to optimize signal-to-noise measurements, it is essential that the electrodes have conformal contact with the human skin. Therefore, we integrated conventional electrodes which use gel while maintaining the features described at the beginning of this report. The E.E.S. will include three electrodes, each in a form of a serpentine like mesh with exposed metal that contacts the skin directly, for recording (REC, electrode1), ground (GND), and reference (REF, electrode2).

The fabrication process begins by spin-coating PDMS on glass slide or Si wafer as a substrate. On top of the UV exposed PDMS, polyamide (PI) must be spin-coated in order to evaporate gold on substrate by electron beam (e-beam) lithography. The gold pattern will be made by using photoresist mask and wet etching, and will be encapsulate with polyamide (PI) to protect gold pattern from any types of defect. This encapsulation is done with the use of a photoresist mask and oxygen plasma etching. After that, the E.E.S. device must be picking up from substrate in order to evaporate chromium-silica through e-beam. It will help to bond device on to the elastomer. The E.E.S. will be laminate on the UV exposed elastomeric substrate (0.5-1.0mm thick) to promote bonding, and water soluble tapes will be remove by applying water. An ACF cable will be attached to the gold pads of E.E.S. and will connects the device to a PCB board.

B.3. S.C.U. (Signal Conditioning Unit)

The Signal Conditioning Unit acts upon the signal that is received from the E.E.S system and with the assistance of its component filters outputs a signal that ranges from 10 to 500 Hertz while possessing a gain of 10000. As shown in the block diagram of the Signal Conditioning Unit, the signal will pass through a variety of components described in detail below.

- Instrumentation Amplifier : Instrumentation amplifier is a type of differential amplifier that eliminates the need for input impedance matching and thus makes the amplifier particularly suitable for use in measurement and test equipment. The low noise of the AD620 allows its use in ECG monitors where high source resistances of 1 MΩ or higher are not uncommon. The AD620’s low power and low supply voltage requirements make it an excellent choice for battery-powered data recorders.

- Negative Feedback : The negative feedback circuit is defined as taking the op-amp's output voltage and coupling it to the inverting input and it is the key to having a self-stabilizing system. This ability to stabilize the system allows the op amp to function in the active region and prevents it from being saturated. Another advantage of using an op amp with negative feedback is that the op amp voltage gain does not matter

- Low-pass Filter : The low pass filter is in charge of attenutating  all the signals above the cut off frequency which is determined by an RC constant. An ideal low pass filter will allow the signals below the cutoff frequency to pass through while attenuating the high frequency signals. The cutoff frequency was calculated using the formula shown below in reference to the figure.

Fig. 4. Typical 2nd order low pass filter. Taken from Reference [10]

The advantage of this configuration as opposed to a simple low pass filter is that the op-amps high input impedance prevents excessive loading on the filters output while its low output impedance prevents the filters cut-off frequency point from being affected by changes in the impedance of the load. Our circuit consists of two 2nd order low pass filters connected together to create a 4th order low pass filter. A second-order passive low pass filter the gain at the cutoff frequency will be equal to 0.7071 x 0.7071 = 0.5Vin (-6dB). The formula that is used to calculate the gain of this filter is G = -R2/R1.

- High-pass Filter : A first-order Active High Pass Filter attenuates low frequencies and passes high frequency signals. The cutoff frequency was calculated using the formula shown below in reference to the figure.

Fig. 5. Typical high pass filter. Taken from Reference [11]

When designing active filter circuits it is important to choose the correct op-amp for the circuit as the loss of high frequency signals may result in signal distortion. The gain of the LME49990 reduces until it reaches unity gain, or its transition frequency which is about 1MHz. This causes the op-amp to have a frequency response curve very similar to that of a first-order low pass filter.

- Band-pass Filter : The Band pass filter is a device that accepts all signals with frequencies within a certain range and rejects signals that are outside this range. The Band pass Filter is normally made by using two low pass filter in conjunction with a high pass filter.

Fig. 6. Typical band pass filter. Taken from Reference [9]

The Bandpass filter on our circuit was made using this method as the high pass filter and the 4th order  low pass filter ,described above, were placed in succesion to create a bandpass filter.

B.4. Convertor (A/D)

In order for the microcontroller to be able to successfully process the intended command of the user, the signal that arrives from the Signal Conditioning Unit needs to be converted into a digital signal for it to be analyzed. The analog to digital converter that we use is the ADS7183. It is a low power, single 16 bit sampling analog-to-digital converter. The result of the converter should be a sequence of digital values that have converted a continuous-time and continuous amplitude analog signal to a discrete time and discrete amplitude signal.

B.5. Micro-Controller

The micro-controller will take the digital from the Analog-to-Digital signal converter and then analyze it using the predetermined commands that were programmed into the PIC. This microcontroller will then bring about the desired output which will influence our device and perform the necessary action on the device that we wish to control. The microcontroller that we use is a PIC. The PIC consists of comparators which will compare the amplitude of the input to the expected amplitude ranges of the signal and it will create the necessary action.

Fig. 7. Conditions where Micro-Controller will act

- Schematic of Overall System

Fig. 8. Overall Schematic of S.C.U.

Fig. 9. Instrumentation Amplifier Schematic

Fig. 10. Negative Feedback

Fig. 11. Band-pass Filter

Fig. 12. Analog to digital Converter to microcontroller

- Simulations & Calculations

- Low-pass Filter: The low pass filter was simulated with a cutoff frequency of around 500 Hz. If we observe the simulation, we will see that there is a constant output voltage below the cutoff frequency. As the frequency passes the cutoff frequency, the output voltage starts to decrease and ultimately reaches a near zero value. The cutoff frequency was calculated using the formula:

Fig. 13. Cutoff Frequency Formula. Taken from Reference[10]

Fig. 14. Simulation wave of Low-pass Filter

- High Pass Filter: The High pass filter that we used in our circuit was calculated to be around 10 Hz. As the frequency of the signal wave remains below the cutoff frequency, the amplitude of the signal will near zero or greater but as the frequency gets closer to the cutoff frequency, the amplitude reaches its peak value and this maximum is maintained. The cutoff frequency was calculated using the following formula

Fig. 15. Cutoff Frequency Formula. Taken from Reference [11]

- Instrumentation Amplifier: The aim of the instrumentation amplifier is to amplify the input signal and produce a signal that consists of a very high gain. The gain of the instrumentation amplifier is determined by a single resistor or Rgain.

The gain of this amplifier can be controlled by a single external resistor. The formula that was used to calculate;

Fig. 16. G Factor & Gain Resistance Formula. Taken from Reference [5]

We arrive at this value of 49.4 K because of the  the internal gain resistors in the instrumentation amplifier are trimmed to 24.7 K each.

Requirements & Verification

| Modules | Requirements | Verification |
| --- | --- | --- |
| E.E.S. | The total thickness of device must be thinner than 2m. | Perform profile measurements using profilometer to verify the total thickness. |
|  | The noise-to-ratio of E.E.S. electrodes must be in ±1mV which is a comparable with commercial EMG sensor. | Setup commercial transmitter and software. Compare signal-to-noise ratio for both commercial and E.E.S. electrodes, and observe the impedance of ±1mV. |
|  | Connection from three electrodes to gold pads of E.E.S. must be made without any disconnections. | Run resistance measurements on probe station. |
|  | Connection from gold pads of E.E.S. to ACF cable must be made without any disconnections. | Run resistance measurements by multimeter. |
| S.C.U. | The output of S.C.U. must have a signal that ranges from 10-500 Hz and has a gain of 10000. | Measure the output voltage of the circuit using a multimeter, and calculate the gain of the circuit after the output voltage is measured. Also, measure the frequency of the signal that is obtained from the signal conditioning unit with the assistance of the multimeter. |
|  | Sub-Requirements | Sub-Verifications |
|  | Instrument Amplifier: The instrument amplifier should increase the gain of the voltage by 10000. | Apply an input signal to the AD620 chip with the assistance of a function generator, and measure output voltage of the AD620. Calculate the gain of the amplifier using the output voltage and the input voltage. G = V0/(V(pin1)-V(pin8)] |
|  | Low Pass Filter: The voltage of the output signal should be constant when the frequency of the signal is less than the cutoff frequency of 500 Hz. | Apply input signal, lower frequency than 500 Hz, and measure the signal at the output of the circuit by the oscilloscope. After observing output voltage, change the input signal to get closer to the cutoff frequency and check output voltage difference. Either the voltage should not change or an output voltage close to zero. |
|  | High Pass Filter: The voltage of the output signal should be constant when the frequency of the signal is greater than the cutoff frequency of 10 Hz. | Apply input signal, greater frequency than 10 Hz, and measure the signal at the output of the circuit by the oscilloscope. After observing output voltage, change the input signal to get closer to the cutoff frequency and check output voltage difference. Either the voltage should not change or an output voltage close to zero. |
| Convertor (A/D) | The analog to digital converter must convert the signals received from the S.C.U. into discrete time and discrete amplitude digital signal. | Apply an analog signal and verify that the ripple from the converter is less than 20mV by oscilloscope. |
| Micro- Controller | The microcontroller must send out the correct signal correspond to the different signals from the actions of user. | Apply a different DC voltage signal to the input pin of the microcontroller and check the output signals are correct correspond to the input signal. |
| Power Supply | The power supply is the main power source for all the components in our system. It will adequately power all the parts of our system | Measure the Power to each of the components powered by the supply with the assistance of the multimeter. |
|  | Sub-Requirements | Sub-Verifications |
|  | Microcontroller Power: The Microcontroller with A/D Convertor must be power by a 5V power supply. | Measure the Vdd by the multimeter so check Vdd pin gets the required voltage of greater than 4.8V |
|  | S.C.U. Power: Each individual chip must have greater than 4.8V. We must ensure that all the op amps and other IC’s get the desired power. |  |

Tolerance Analysis

The E.E.S. electrodes must work to a tight tolerance since accurate EMG signals must be detected compare to conventional EMG sensor. If we get far different signals from E.E.S., the device is not actually performing as an EMG sensor. With the assistance of an oscilloscope and documented signals from previous EMG tests, we will be able to determine if the signal from the E.E.S. is as accurate as the signals from the EMG measured by conventional electrodes. As long as the signal to noise ratio is within ±1mV, our E.E.S. electrodes are comparable with conventional EMG sensor.

- Since our goal is to use E.E.S. electrodes as a Human-Machine Interface, two sets of the E.E.S. electrodes will be mounted on left and right forearm. The EMGs of both electrodes will be measured with four different actions: bending the hands to the left, bending the hands to the right, rotating both hands in a circle, and squeezing both fists. The signals must be classified enough to use it as different command to control the quad-rotor. If four different actions are not giving enough different signals compare to others, than we must consider changing the location of E.E.S. electrodes or using different actions beside these four actions. Furthermore, using additional E.E.S. electrodes on other parts of the body can solve this problem, giving expanded available range of motions and classifications of movement.

Ethical Issues

We adhere to the statements of the IEEE Code of Ethics that pertain to our project as follows:

- “to be honest and realistic in stating claims or estimates based on available data”

We will ensure that all calculations are accurate. All the conclusions drawn from experimental procedures will be supported by data calculations and simulations.

- “To seek, accept, and offer honest criticism of technical work, to acknowledge and correct errors, and to credit properly the contributions of others.”

Since this is a group project, each member of the team will assist with the design and provide feedback that will help improve other member’s designs and the overall design of our circuit.

All material and information that was used will be appropriately credited in the References section.

- “to improve the understanding of technology; it’s appropriate application, and potential consequences”

We will apply the skills that we have acquired and learned thus far and implement these skills to the best of our abilities and create an exciting E.E.S human interface design.

Safety

In order to make a worthwhile EES human interface system, we will need to guarantee that the product will be safe for consumer use. When creating the S.C.U., we use three resistors across each electrode that interacts with the human skin to serve as protection for the user, this will prevent backflow of voltage and shock the user. This will comply with the safety of the humans and the device.

In order to guarantee circuit and human safety, the sprinkler robot will use proper grounding. The grounding will serve as a means for safety because it will prevent shocks to the user in case the electrical insulation fails.

In order to be sure that our project is clear of hazardous materials, all the parts used in our circuit are compliant with the Restriction of Hazardous Substances (RoHS) Directive. This implies that the entire circuit is free of lead, mercury, cadmium, hexavalent chromium, Polybrominated biphenyls, and Polybrominated diphenyl ether.

Cost & Schedule

- Cost Analysis

*We did not include any supported equipment and materials for E.E.S. in cost chart

- Labor

|  | $/Hour | Overhead | Total hours | Subtotal ($) |
| --- | --- | --- | --- | --- |
| Woosik Lee | 50 | 125 | 144 | 18000 |
| Ohjin Kwon | 50 | 125 | 144 | 18000 |
| Nithin Reddy | 50 | 125 | 144 | 18000 |
| Note : Total hours = 12hous x 12weeks |

- Parts

| Description | Component | # | Manufacturer | Vendor | Cost/Unit ($) | Total Cost |
| --- | --- | --- | --- | --- | --- | --- |
| Negative Feedback |
| Op Amp (OP97FSZ) | OP97 | 1 | Analog Devices | Newark | 2.67 | 2.67 |
| Capacitor (C1206C103F3GACTU) | 0.01uF | 1 | Kemet | Mouser | 1.95 | 1.95 |
| Register (TNPW120610K0BEEA) | 10K Ohms | 1 | Vishay |  | 0.55 | 0.55 |
| Register (ERA-2AEB113X) | 11K Ohms | 1 | Panasonic Electronic Components |  | 0.6 | 0.6 |
| Register (CRT1206-DY-2203ELF) | 220K Ohms | 1 | Bourns |  | 0.1 | 0.1 |
| Instrumentation Amplifier |
| Instrument Amp (AD620ARZ) | AD620 | 1 | Analog Devices | Newark | 7.24 | 7.24 |
| Op Amp (UA741IDT) | 741 | 2 | STMicroelectronics | Mouser | 0.6 | 1.2 |
| Capacitor (GRM31C5C1H104FA01L) | 0.1uF | 1 | Murata |  | 1.67 | 1.67 |
| Register (MCW0406MD2001BP100) | 2K Ohms | 1 | Vishay/BC Components |  | 0.94 | 0.94 |
| Register (ERJ-8ENF2202V) | 22K Ohms | 2 | Panasonic Electronic Components |  | 0.1 | 0.2 |
| Register (TNPW12061M00BEEA) | 1M Ohms | 1 | Vishay |  | 1.11 | 1.11 |
| Bandpass Filter |
| Op Amp (LME49990MA/NOPB) | L49990 | 3 | Texas Instruments | Mouser | 2.58 | 7.74 |
| Potentiometer (M64Y503KB40) | P-50K | 1 | Vishay/Spectrol |  | 1.78 | 1.78 |
| Potentiometer (PV36W105C01B00) | POT-1M | 1 | Murata |  | 1.67 | 1.67 |
| Capacitor (GRM31C5C1H104FA01L) | 0.1uF | 2 | Murata |  | 1.67 | 2.34 |
| Capacitor (ECH-U1H154GX9) | 0.15uF | 1 | Panasonic Electronic Components |  | 1.44 | 1.44 |
| Capacitor (ECW-U1C274JC9) | 0.27uF | 1 | Panasonic Electronic Components |  | 1.09 | 1.09 |
| Capacitor (ECW-U1C334JC9) | 0.33uF | 1 | Panasonic Electronic Components |  | 1.13 | 1.13 |
| Capacitor (C1210C475J3NACTU) | 4.7uF | 1 | Kemet |  | 0.85 | 0.85 |
| Capacitor (C1210C106J8NACTU) | 10uF | 1 | Kemet |  | 1.09 | 1.09 |
| Register (RP73PF2A174RBTDF) | 174 Ohms | 1 | Teconnectivity | Newark | 0.35 | 0.35 |
| Register (RG3216P-2000-B-T1) | 200 Ohms | 1 | SUSUMU |  | 0.797 | 0.797 |
| Register (CRCW1206422RFKEA) | 422 Ohms | 1 | Vishaydale |  | 0.1 | 0.1 |
| Register (RP73D2B475RBTG) | 475 Ohms | 1 | TEconnectivity/HOLSWORTHY |  | 1.27 | 1.27 |
| Register (ERA-8AEB1241V) | 1.24K Ohms | 1 | Panasonic |  | 0.328 | 0.328 |
| Register (RP73D2A2K94BTG) | 2.94K Ohms | 1 | Teconnectivity/HOLSWORTHY |  | 1.35 | 1.35 |
| PIC16F88 |  | 2 | Microchip |  | 4.10 | 8.20 |
| Total Cost of Parts: | 47.20 |

- Grand Total Cost

| Labor | Parts | Total |
| --- | --- | --- |
| $54,000 | $47. 20 | $54,047.20 |

- Schedule

| Week | Task Description | Member |
| --- | --- | --- |
| 2/6 | Proposals Due |
|  | Research how to design E.E.S. | Woosik Lee |
|  | Update Proposal | Ohjin Kwon |
|  | Research how to design microcontroller | Nithin Reddy |
| 2/11 | Mock Design Review Due |
|  | Start to design E.E.S. | Woosik Lee |
|  | Start to design S.C.U. | Ohjin Kwon |
|  | Start to design microcontroller | Nithin Reddy |
| 2/18 | Finish to design E.E.S. | Woosik Lee |
|  | Finish to design S.C.U. | Ohjin Kwon |
|  | Finish to design microcontroller & Search to order parts | Nithin Reddy |
| 2/25 | Design Review Due |
|  | Mechanical testing on E.E.S. design & Design review report | Woosik Lee |
|  | Finish to design S.C.U. & Design review report | Ohjin Kwon |
|  | Start to design PCB board & Design review report | Nithin Reddy |
| 3/4 | Modify E.E.S. & Start fabricating E.E.S. devices | Woosik Lee |
|  | Modify S.C.U. | Ohjin Kwon |
|  | Modify microcontroller | Nithin Reddy |
| 3/11 | Individual Progress Reports Due |
|  | Continue fabricating E.E.S. & Run tests | Woosik Lee |
|  | Finish to design PCB board & Submit board to manufacture | Ohjin Kwon |
|  | Assemble the boards | Nithin Reddy |
| 3/18 | Spring Break |
| 3/25 | Mock-up Demos Due |
|  | Mock-up Demos (E.E.S.) | Woosik Lee |
|  | Mock-up Demos (S.C.U.) | Ohjin Kwon |
|  | Mock-up Demos (Microcontroller) | Nithin Reddy |
| 4/1 | Mock-up Presentation Due |
|  | Mock-up Presentation (E.E.S.) | Woosik Lee |
|  | Mock-up Presentation (S.C.U.) | Ohjin Kwon |
|  | Mock-up Presentation (Microcontroller) | Nithin Reddy |
| 4/8 | Testing E.E.S. | Woosik Lee |
|  | Testing S.C.U. | Ohjin Kwon |
|  | Testing Microcontroller | Nithin Reddy |
| 4/15 | Modify E.E.S. & Verifications | Woosik Lee |
|  | Modify S.C.U. & Verifications | Ohjin Kwon |
|  | Modify Microcontroller & Verifications | Nithin Reddy |
| 4/22 | Demos & Presentation Due |
|  | Demo, Final testing E.E.S. & Assemble everything | Woosik Lee |
|  | Demo, Final testing S.C.U. & Assemble everything | Ohjin Kwon |
|  | Demo, Final testing Microcontroller & Assemble everything | Nithin Reddy |
| 4/29 | Presentation, Final Papers, Lab Notebooks, Checkout Due |
|  | Final Paper (Testing & Verification) | Woosik Lee |
|  | Final Paper (Design Procedure) | Ohjin Kwon |
|  | Presentation, Final Paper (Introduction & Conclusion) | Nithin Reddy |

Reference

- Dr. Scott Day. Important Factors in Surface EMG Measurement. Bortec Biomedical Ltd.

- Hermie J. Hermens, Bart Freriks, Catherine Disselhorst-Klug, Gu¨nter Rau. Development of recommendations for SEMG sensors and sensor placement procedures. Roessingh Research and Development.

- Carlo J.DeLuca, MikhailKuznetsov, L.DonaldGilmore, SergeH.Roy. Inter-electrode spacing of surface EMG sensors: Reduction of crosstalk contamination during voluntary contractions. NeuroMuscular Research Center, Boston University.

- Carlo J. De Luca. Surface electromyography: Detection and Recording. DelSys Incorporated, 2002.

- Instrumentation Amplifier AD620, datasheet, Analog Devices, Inc, 2004. Available at: http://www.analog.com/en/specialty-amplifiers/instrumentation-amplifiers/ad620/products/product.html

- Operational Amplifier OP97, datasheet, Analog Devices, Inc, 1997-2009. Available at: http://www.analog.com/en/all-operational-amplifiers-op-amps/operational-amplifiers-op-amps/op97/products/product.html

- E-Series Ultra-low Distortion, Ultra-low Noise Operational Amplifier LME49990, datasheet, Texas Instruments, 2011. Available at: http://www.ti.com/product/lme49990

- General Purpose Single Operational Amplifier UA741, datasheet, STMicroelectronics, 2001. Available at: http://www.datasheetcatalog.com/datasheets_pdf/U/A/7/4/UA741.shtml

- Instrumentation Amplifier, Wikipedia web page. Available at: http://en.wikipedia.org/wiki/Instrumentation_amplifier

- Low Pass Filter, Wikipedia web page. Available at: http://en.wikipedia.org/wiki/Low-pass_filter

- High Pass Filter, Wikipedia web page. Available at: http://en.wikipedia.org/wiki/High-pass_filter